## Slide 1: CARDIAC CATHETERS

## Slide 2: HISTORY
First cardiac catheterization was done
In 1929
Werner Forssmann
Rubber catheter
Through his own antecubital vein up to pulm artery

## Slide 3: CORONARY CATHETERS
Hollow flexible tube that can be inserted through a narrow opening into the body cavity,duct or vessel to diagnose and treat heart conditions.

## Slide 4: BASIC FUNCTIONS
Inject contrast agents
To provide back up support to pass coronary balloon /stent into coronary artery against resistance and across target stenosis
To monitor systemic arterial pressure

## Slide 5: IDEAL CHARACTERISTICS
Better torque control
Pushability
Flexibility
Trackability
Radiopacity
Atraumatic tip
Kink resistance

## Slide 6: CATHETER MATERIALS
Angiographic catheters are made from synthetic and semi synthetic polymers
Dacron – Maneuverable , flexible
Nylon – increase bursting pressure
PVC – increased thrombogenecity ,soft and flexible
Polyethylene
Fluoropolymers
Polyurethane
Silicon –most common choice; inactive to body fluids
Radiopacity by incorporating Ba, Bi, Ir

## Slide 7: SIDE HOLE
Useful where pressure gets frequently damped as in RCA interventions ,CTO interventions or sole surviving artery or LMCA interventions.
Advantages
Prevent catheter damping
Allow additional blood flow out of the tip to perfuse the artery
Avoid catastrophic dissections in the ostium of the artery

## Slide 8
Disadvantages
Incorrect pressure tracing (Aortic pressure ;not the coronary pressure is being monitored)
Suboptimal opacification
Reduction of back up support
Kinking at side hole

## Slide 9: BACK UP SUPPORT
Ability of guide catheter to remain in position
Provide a stable platform for the advancement of interventional equipments.
3 major types of back up support : Active, Passive & Balanced
i)Active support
Support provided by the operator’s manipulation
ii) Passive support
Relies on the catheter shape and engagement in the artery
iii)Balanced support
Aim to achieve a balance between active and passive support offering a combination of both features

## Slide 10: CATHETER SIZE
Catheters are named by their outer diameter in French (Fr)
1 Fr = 0.33 mm
5 – 10 Fr available
Internal diameter is represented in inches
Length in cm – 80, 100,110 (125 cm for very tall )
6F 7F 8F

## Slide 11: CORONARY ANATOMY VARIATION
Coronary ostial location
i) High
ii) Low
iii) Anterior
iv) Posterior
Coronary ostial orientation
i) Superior
ii) Horizontal
iii) Inferior
iv) Shepherd crook RCA

## Slide 12: TYPES OF CATHETERS
There are 2 types of catheters;
1) Diagnostic catheters
2) Guiding catheters
DIAGNOSTIC VS GUIDING CATHETERS
Diagnostic catheters | Guiding catheters
1. Thicker shaft | 1. Thin shaft
2. Smaller internal diameter | 2. Larger internal diameter
3. Tapering tip | 3. Non tapering tip
4. Two layered construction | 4. 3 layered construction

## Slide 13: DIAGNOSTIC CATHETERS
Used to investigate and evaluate patients with suspected or known CAD
2 layered construction ;
i) Outer layer – non-thrombogenic and lubricious
ii) Inner layer - lubricious

## Slide 14: GUIDING CATHETERS
Primarily used to deliver interventional equipments (GW, balloon ,stent) during coronary intervention
3 layered construction ;
i) Outer layer – lubricious and non thrombogenic
ii) Middle layer – braided (eg. Stainless steel)
iii) Inner layer – lubricious (eg. Teflon)

## Slide 15
LARGE GC | SMALL GC
High bleeding risk | Lower bleeding risk
Greater coronary opacification | Less coronary opacification
More passive support | Less passive support
More complex PCI possible | Less complex PCI possible
Better torque transmission | Poor torque transmission

## Slide 16: DIFFERENT VARIETIES OF CATHETERS
RIGHT HEART CATHETERS | ANGIOGRAPHIC CATHETERS | PREFORMED CATHETERS
Cournand Lehman Goodale-lubin | Gensini NIH Eppendorf Pigtail | Judkins Amplatz Schoonmaker MP EBU Coronary bypass catheter, etc
CATHETER CHOICE FOR CORONARIES
LMCA | RCA | Universal | Universal Radial | Graft and IMA
JL AL EBU | JR AR 3DRC | MP Sones | Tiger Jacky | RCB LCB IMA

## Slide 17
TIGER & JACKY CATHETER
Universal radial catheter
Both LMCA & RCA can be cannulated
Has an end hole
High flexibility
Size : 5 Fr & 6Fr available
Tiger - one SH & Jacky – 2 SH

## Slide 18: JUDKINS CATHETER
Designed by Melvin Judkins
Made of polyethylene
Designed for femoral route
Distance between primary curve and secondary curve may vary (3,3.5,4,4.5)
Has an end hole with or without SH
2 types ; Judkins Right (JR) & Judkins Left (JL)

## Slide 19: EXTRA BACK UP (EBU) CATHETER
Designed for LMCA cannulation
2 curves; primary and a long secondary curve
Long 2 curve provides stable platform
Better support and stability
Coaxial alignment
Nylon shaft
Size : 6 – 8 Fr & 100 cm

## Slide 20: AMPLATZ CATHETER
Designed by Kurt Amplatz
Provide good back up support
Best in case of : short LM with downgoing LCX
shepherd crook RCA
Calcified lesions , CTO s
Disadvantage : increased chances of dissection
2 types ; Amplatz Left ( AL) & Amplatz Right (AR)
AR | AL

## Slide 21
A speciality of amplatz catheter is that if you want to disengage the catheter from the coronary ostium, you have to push the catheter rather than withdrawing it
If you try to withdraw – gets more deeply engaged

## Slide 22: 3DRC ( 3 Dimentional right curve)
For tortious ,bent anatomy and posterior / superior take-off of RCA.

## Slide 23: CORONARY BYPASS CATHETERS
RCB CATHETER
For right coronary venous bypass graft
Resembles JR 4
LCB CATHETER
For left coronary venous bypass graft
IMA CATHETER(internal mammary artery catheter)
Resembles JR
Thicker primary curve (80 and longer tip

## Slide 24
COURNAND CATHETER
Designer – Andre Cournand in 1939
End hole , no side hole
Made of Woven Dacron with outer coating of polyurethane.
Has a very gradual distal curve, tapered tip.
Use - all purpose right heart catheter
- wedge pressure measurements(no side hole)
Size – 5 to 8 Fr & length 100 and 125 cm

## Slide 25: SCHOONMAKER MULTIPURPOSE
Designer – Dr Fred W Schoonmaker
Construction – polyurethane with inner wire braid
Single curve with straight tip
Has end hole
Use – CAG :both nature vessel and graft
- RHC
- Ventriculography
2 types : MPA( 120 curve) & MPB(90) curve
MPA 1 | MPA 2 | MPB 1 | MPB 2
1 EH only | 1 EH & 2 SH | 1 EH only | 1 EH & 2 SH

## Slide 26: NIH CATHETER
Has only side holes and no end holes
3 pairs of oppositely directed side holes
Made of woven dacron with nylon reinforcement
Diameter – 4to 8 Fr
Length – 50 to 125 cm
Uses : - Ventriculography

## Slide 27: PIGTAIL CATHETER
Designer – Judkins
Made of woven dacron coated with polyurethane or polyethylene
Has an end hole and 3to 6 pairs of side holes in the terminal 5 cm
Tapered tip ,terminal 5 cm is coiled in a tight loop
Size : 4 to 7 Fr and 65 to 110 cm in length
Uses : LV , RV angiography
Aortography
Pulmonary angiography
Pericardiocentesis
Pressure monitoring

## Slide 28
Advantages : atraumatic
Disadvantage : thrombogenecity and not suitable for prolonged hemodynamic monitoring.
HALO CATHETER
5 Fr catheter
Perpendicular helical tip which is inwardly and upwardly directed
Side holes – located on the helix
Use : Electrophysiological study

## Slide 29: GOODALE LUBIN CATHETER
Bird’s eye catheter
2 oval shaped laterally opposed SH near EH
Made of woven dacron and coated with polyurethane
Size : 4 – 8 Fr & 80 – 125 cm
Use: Rt heart pressure measurement
Variation : standard wall – Cournand
thin wall – Lehman

## Slide 30: LEHMAN CATHETER
Thin wall variation of Cournand
Tapered closed tip
Multiple SH and no EH
Short distal curve
Size : 4 – 9 Fr & 50 – 125 cm
Use : ventriculography

## Slide 31: GENSINI CATHETER
45 smooth curve , tapered tip
Has an end hole and 3 laterally opposed oval SH
Made of woven dacron ,coated with polyurethane
Use :right and left angiographic studies
Disadvantage : more arrhythmogenic

## Slide 32: EPPENDORF CATHETER
Have a gentle curve distally
Made of woven dacron and coated with polyurethane
20 cm proximal to hub is reinforced with nylon
Has 6 laterally opposed side holes and no end hole
Size : 7 – 8 Fr & 100 – 125 cm
Less stiff and more torque control
Use : visualising RV, LV,arterial ,pulmonary vasculature
ventriculography
aortography

## Slide 33: SONES CATHETER
Made of woven dacron
Covered by polyurethane
Long 45 distal curve and tapers near the tip
Has an end hole & a pair of side holes
Used for left ventriculography via brachial approach

## Slide 34: MCQ s
Which among the following is a catheter without an end hole?
a) Tiger
b) Judkins right
c) Goodale lubin
d) NIH catheter

## Slide 35
2 . Which among the following is also known as bird’s eye catheter?
a) Eppendorf
b) NIH catheter
c) Goodale lubin
d) Amplatz Left

## Slide 36
3. Which among the following catheter is not used for RCA cannulation?
a) EBU
b) JR
c) Tiger
d) AR